## Slide 1: Color characterization of virtual reality devices using professional- and consumer-grade instruments
Wei-Chung Cheng
Division of Imaging, Diagnostics, and Software Reliability
Office of Science and Engineering Laboratories
Center for Devices and Radiological Health
U.S. Food and Drug Administration
*The mention of commercial products, their sources, or their use in connection with material reported herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services. This is a contribution of the US Food and Drug Administration and is not subject to copyright.

## Slide 2: Introduction
Color characterization is important when AR/VR devices are used for research
  Vision research, color science, lighting, automotive, behavior science, and cognitive science
Missing standard testing method for color characterization for AR/VR devices
  Information Display Measurements Standard (IDMS)
  International Committee for Display Metrology (ICDM) recommends many test methods for spatial resolution; not much update on color measurement
  Costly instrumentation and professional set-up are required; not practical for end users to conduct QA/QC remotely, especially during the pandemic
Research question: “Can consumer-grade instruments be used to measure color characteristics for VR devices?”

## Slide 3: Overview
Color characteristics of a VR headset
Professional-grade instrument
Measurement dataset
Consumer-grade instrument
Error analysis
Calibration method
VR
Professional
Consumer
Error
Truth
Calibration
Output

## Slide 4: Professional-grade Instrument
Spectroradiometer (PR-730, Photo Research)
  380-780 nm @ 1 nm, 75 mm lens, 2֯ measuring aperture
  $25,000
Telescopic measurement method
Newtonian viewing configuration
  No optics between headset and instrument
Criteria*
  Light source uniformly fills the measurement field
  Angular aperture is uniformly illuminated and fills the entrance pupil
  *Penczek, JSID, 25(4), 2017

## Slide 5: Consumer-grade Instrument
Spyder X Elite, DataColor
  Color calibration kit (CCK) intended for generating ICC color profiles for computer displays
  Colorimeter
  $250
Contact measurement method
  Sensor placed on ocular lens directly

## Slide 6: Use CCK as a Colorimeter
Spyder X Elite Software
  “Colorimeter”
  CIE xyY => CIELAB
  Backlighting options
    “Wide LED”
    “Standard LED”
    “General”
    “GB LED”
  Heuristic of LED spectrum to improve accuracy
  No documentation

## Slide 7: VR Headset
OLED-based VR headset (Rift, Oculus)
Driven through HDMI by Matlab on Windows 10
External display of 2160x1200 pixels
No rendering software included in the experiments
sRGB ICC color profile
Each color patch was generated as a full-screen test pattern

## Slide 8: Setup
VR
Spectroradiometer
CCK Sensor
Spectrum
CIE xyY
CIELAB
CIELAB
ΔE
CCK S/W
Matlab
Matlab
Win10
Patch
Calibration

## Slide 9: Datasets
 | Professional-grade | Consumer-grade
Primary colors (3 patches) | ✓ | 
Color cube (602 patches) | ✓ | 
ColorChecker (24 patches) | ✓ | ✓
Datasets #1 and #2: report color gamut of Rift
Dataset #3: compare professional-grade with consumer-grade instruments

## Slide 10: Measurement Data #1 (3 patches)
Spectra of primaries
  Narrow-band spectra of LEDs
  some overlapping between R/G, G/B
Chromaticity gamut area
  IDMS 5.18
Standard color spaces
  sRGB < DCI-P3 < AdobeRGB < Rec. 2020
Doesn’t represent the darker shades

## Slide 11: Measurement Data #2 (602 patches)
Color gamut volume
  IDMS 5.32
  6 slides per volume
  11x11 grid per side (121 points)
  602 points in total
Calculate volume by counting ΔE3
  Rift: 1,273,100 ΔE3
  sRGB: 830,732 ΔE3
  53.25% larger
  Calculated by IDMS Matlab code

## Slide 12: Measurement Data #3 (24 patches)
Measured spectra/shades
ColorChecker
  RGB (1 spike)
  CYM (2 spikes)
  Grayscale (3 spikes)

## Slide 13: Measurement Data #3 from Consumer-grade
Color errors calculated in CIEDE2000
  “General” backlight
  Mean: 1.02
  Maximum: 3.21
  Standard deviation: 0.68
Larger color errors
  Patches #1, #4, #7, #15, #23 and #24

## Slide 14: CCK Accuracy vs Backlight Mode
“Backlight modes”
  The mean errors are about the same

## Slide 15: Error Analysis
Fit CIEXYZ values with linear regression
  R2 > 0.99 => X, Y, and Z have perfect linearity
CIE Z deviated
  Error increases as Z decreases
  Larger errors for darker shades

## Slide 16: Calibration
Scale coefficients of X and Z to equalize their slopes with Y

## Slide 17: Calibration Results
Before calibration
  1.02 ΔE00 (“General”)
After calibration
  0.57 ΔE00 (“General”) or 44% reduction
  “Standard LED” has the lowest ΔE00

## Slide 18: Discussion
Rendering software is important but not included in the experiments
Consumer-grade instruments may be sufficient for measuring VR colors but not for photometry such as luminance
Using a CCK as a colorimeter is a useful feature, but not all CCKs support it
CCK accuracy depends on color shades (e.g., pink shades in digital pathology)
IDMS Appendix A7 provides calibration methods for telescopic meters, but not for contact-type

## Slide 19: Conclusions
Professional-grade instruments can be used by researchers to measure VR headsets in a lab setting.
Consumer-grade instruments can be used by end-users to conduct on-site quality assurance/quality control procedures.
The proposed calibration method can be used to improve accuracy of color calibration kits.
This study may help end-users make informed decisions on adopting consumer-grade instruments that are much cheaper than professional-grade instruments.

## Slide 20: Medical Extended Reality Program
https://www.fda.gov/medical-devices/medical-device-regulatory-science-research-programs-conducted-osel/medical-extended-reality-program-research-medical-extended-reality-based-medical-devices